## Modbus
| Module address: | 16
Modbus RTU | COM-port:
| Port number: | 1 (X21)
| Baud Rate: | 38400 (9600...38400)
| Data bytes: | 8
| Stop bits: | 1
| Parity: | None
| RTS-Control: | DISABLE
| FloatFormat: | 3 - FP LB(2,1,4,3)
COM-port:
| Port number: | 2 (X22)
| Baud Rate: | 38400 (9600...38400)
| Data bytes: | 8
| Stop bits: | 1
| Parity: | None
| RTS-Control: | DISABLE
| FloatFormat: | 3 - FP LB(2,1,4,3)
Modbus TCP | TCP-port:
| Port number: | 502
| FloatFormat: | 3 - FP LB(2,1,4,3)
Ethernet, TCP port: 502 (can be reconfigured as all other physical links parameters above)
| Method | ISO 6976
| Function Code | 3
Address: | Name | Type | Data type
0 | NumMet (Method number) | System Settings | UINT16
1 | Num (Number Analysis) | Current analysis Parameters | UINT16
2 | Day (Day of the month of analysis) | Current analysis Parameters | UINT16
3 | Month (Month of analysis) | Current analysis Parameters | UINT16
4 | Year (Year of analysis) | Current analysis Parameters | UINT16
5 | Hour (Analysis start hour) | Current analysis Parameters | UINT16
6 | Min (Analysis start minute) | Current analysis Parameters | UINT16
7 | Sec (Analysis start second) | Current analysis Parameters | UINT16
8 | FlagErr (Analysis errors) | Current analysis Parameters | UINT16
9 | CRC0 (hash-code calculation unit) | Current analysis Parameters | UINT16
10 | dConc (The final concentration) | Methane | FLOAT
12 | dConc (The final concentration) | Ethane | FLOAT
14 | dConc (The final concentration) | Propane | FLOAT
16 | dConc (The final concentration) | i-Butane | FLOAT
18 | dConc (The final concentration) | n-Butane | FLOAT
20 | dConc (The final concentration) | neo-Pentane | FLOAT
22 | dConc (The final concentration) | i-Pentane | FLOAT
24 | dConc (The final concentration) | n-Pentane | FLOAT
26 | dConc (The final concentration) | n-Hexane+ | FLOAT
28 | dConc (The final concentration) | Carbon dioxide | FLOAT
30 | dConc (The final concentration) | Nitrogen | FLOAT
32 | dConc (The final concentration) | Oxygen | FLOAT
34 | dConc (The final concentration) | Hydrogen | FLOAT
36 | dConc (The final concentration) | Helium | FLOAT
38 | dConc (The final concentration) | Water | FLOAT
40 | ParM33 (Q Sup.vol.(15,15), MJ/m3) | Method parameters | FLOAT
42 | ParM36 (Q Inf.vol.(15,15), MJ/m3) | Method parameters | FLOAT
44 | ParM39 (W Sup.(15,15), MJ/m3) | Method parameters | FLOAT
46 | ParM42 (W Inf.(15,15), MJ/m3) | Method parameters | FLOAT
48 | ParM49 (Q Sup.vol.(15,15),Kwh/m3) | Method parameters | FLOAT
50 | ParM46 (Q Inf.vol.(15,15),Kwh/m3) | Method parameters | FLOAT
52 | ParM52 (W Sup.(15,15), Kwh/m3) | Method parameters | FLOAT
54 | ParM56 (W Inf.(15,15), Kwh/m3) | Method parameters | FLOAT
56 | ParM35 (Q Sup.vol.(15,25), MJ/m3) | Method parameters | FLOAT
58 | ParM38 (Q Inf.vol.(15,25), MJ/m3) | Method parameters | FLOAT
60 | ParM41 (W Sup.(15,25), MJ/m3) | Method parameters | FLOAT
62 | ParM44 (W Inf.(15,25), MJ/m3) | Method parameters | FLOAT
64 | ParM51 (Q Sup.vol.(15,25),Kwh/m3) | Method parameters | FLOAT
66 | ParM48 (Q Inf.vol.(15,25),Kwh/m3) | Method parameters | FLOAT
68 | ParM54 (W Sup.(15,25), Kwh/m3) | Method parameters | FLOAT
70 | ParM58 (W Inf.(15,25), Kwh/m3) | Method parameters | FLOAT
72 | ParM60 (Abs.density(15), kgr/m3) | Method parameters | FLOAT
74 | ParM62 (Rel.density (15)) | Method parameters | FLOAT
76 | ParM64 (Compression factor (15)) | Method parameters | FLOAT
78 | ParM34 (Q Sup.vol.(0,25), MJ/m3) | Method parameters | FLOAT
80 | ParM37 (Q Inf.vol.(0,25), MJ/m3) | Method parameters | FLOAT
82 | ParM40 (W Sup.(0,25), MJ/m3) | Method parameters | FLOAT
84 | ParM43 (W Inf.(0,25), MJ/m3) | Method parameters | FLOAT
86 | ParM50 (Q Sup.vol.(0,25),Kwh/m3) | Method parameters | FLOAT
88 | ParM47 (Q Inf.vol.(0,25),Kwh/m3) | Method parameters | FLOAT
90 | ParM53 (W Sup.(0,25), Kwh/m3) | Method parameters | FLOAT
92 | ParM57 (W Inf.(0,25), Kwh/m3) | Method parameters | FLOAT
94 | ParM61 (Abs.density(0), kgr/m3) | Method parameters | FLOAT
96 | ParM63 (Rel.density (0)) | Method parameters | FLOAT
98 | ParM65 (Compression factor (0)) | Method parameters | FLOAT
100 | ParM5 (Q Sup.vol.(20,25), MJ/m3) | Method parameters | FLOAT
102 | ParM4 (Q Inf.vol.(20,25), MJ/m3) | Method parameters | FLOAT
104 | ParM10 (W Sup.(20,25), MJ/m3) | Method parameters | FLOAT
106 | ParM11 (W Inf.(20,25), MJ/m3) | Method parameters | FLOAT
108 | ParM45 (Q Inf.vol.(20,25),KWh/m3) | Method parameters | FLOAT
110 | ParM55 (W Sup.(20,25), Kwh/m3) | Method parameters | FLOAT
112 | ParM59 (W Inf.(20,25), Kwh/m3) | Method parameters | FLOAT
114 | ParM8 (Abs.density(20), kgr/m3) | Method parameters | FLOAT
116 | ParM9 (Rel.density (20)) | Method parameters | FLOAT
118 | ParM2 (Compression factor (20)) | Method parameters | FLOAT
120 | ParM7 (Q Sup.mol.(25) kJ/mole) | Method parameters | FLOAT
122 | ParM6 (Q Inf.mol.(25) kJ/mole) | Method parameters | FLOAT
124 | ParM14 (Q Sup.mass.(25), MJ/kgr) | Method parameters | FLOAT
126 | ParM13 (Q Inf.mass.(25), MJ/kgr) | Method parameters | FLOAT
128 | ParM3 (Gas molecular weight) | Method parameters | FLOAT
130 | FirstNum (Number of the first analysis of the sample) | Request -Analysis Parameters | UINT16
131 | FlagErr (Analysis error) | Request -Analysis Parameters | UINT16
132 | CRC0 (hash-code calculation unit) | Request -Analysis Parameters | UINT16
133 | NumKal1 (Test gas pasport number of Analyzer 1) | Request -Analysis Parameters | UINT16
134 | Count (Count of analyses of the sample) | Request -Analysis Parameters | UINT16
135 | Status (Response status of the analysis request) | Request -Analysis Parameters | UINT16
136 | FirstDay (Day of the month of the first analysis of the sample) | Request -Analysis Parameters | UINT16
137 | FirstMonth (Month of the first analysis of the sample) | Request -Analysis Parameters | UINT16
138 | FirstYear (Year of the first analysis of the sample) | Request -Analysis Parameters | UINT16
139 | FirstHour (Start hour of the first analysis of the sample) | Request -Analysis Parameters | UINT16
140 | FirstMin (Start minute of the first analysis of the sample) | Request -Analysis Parameters | UINT16
141 | FirstSec (Start Second of the first analysis of the sample) | Request -Analysis Parameters | UINT16
142 | dConc (The final concentration) | Request -Methane | FLOAT
144 | dConc (The final concentration) | Request -Ethane | FLOAT
146 | dConc (The final concentration) | Request -Propane | FLOAT
148 | dConc (The final concentration) | Request -i-Butane | FLOAT
150 | dConc (The final concentration) | Request -n-Butane | FLOAT
152 | dConc (The final concentration) | Request -neo-Pentane | FLOAT
154 | dConc (The final concentration) | Request -i-Pentane | FLOAT
156 | dConc (The final concentration) | Request -n-Pentane | FLOAT
158 | dConc (The final concentration) | Request -n-Hexane+ | FLOAT
160 | dConc (The final concentration) | Request -Carbon dioxide | FLOAT
162 | dConc (The final concentration) | Request -Nitrogen | FLOAT
164 | dConc (The final concentration) | Request -Oxygen | FLOAT
166 | dConc (The final concentration) | Request -Hydrogen | FLOAT
168 | dConc (The final concentration) | Request -Helium | FLOAT
170 | dConc (The final concentration) | Request -Water | FLOAT
172 | ParM33 (Q Sup.vol.(15,15), MJ/m3) | Request -Method parameters | FLOAT
174 | ParM36 (Q Inf.vol.(15,15), MJ/m3) | Request -Method parameters | FLOAT
176 | ParM39 (W Sup.(15,15), MJ/m3) | Request -Method parameters | FLOAT
178 | ParM42 (W Inf.(15,15), MJ/m3) | Request -Method parameters | FLOAT
180 | ParM49 (Q Sup.vol.(15,15),Kwh/m3) | Request -Method parameters | FLOAT
182 | ParM46 (Q Inf.vol.(15,15),Kwh/m3) | Request -Method parameters | FLOAT
184 | ParM52 (W Sup.(15,15), Kwh/m3) | Request -Method parameters | FLOAT
186 | ParM56 (W Inf.(15,15), Kwh/m3) | Request -Method parameters | FLOAT
188 | ParM35 (Q Sup.vol.(15,25), MJ/m3) | Request -Method parameters | FLOAT
190 | ParM38 (Q Inf.vol.(15,25), MJ/m3) | Request -Method parameters | FLOAT
192 | ParM41 (W Sup.(15,25), MJ/m3) | Request -Method parameters | FLOAT
194 | ParM44 (W Inf.(15,25), MJ/m3) | Request -Method parameters | FLOAT
196 | ParM51 (Q Sup.vol.(15,25),Kwh/m3) | Request -Method parameters | FLOAT
198 | ParM48 (Q Inf.vol.(15,25),Kwh/m3) | Request -Method parameters | FLOAT
200 | ParM54 (W Sup.(15,25), Kwh/m3) | Request -Method parameters | FLOAT
202 | ParM58 (W Inf.(15,25), Kwh/m3) | Request -Method parameters | FLOAT
204 | ParM60 (Abs.density(15), kgr/m3) | Request -Method parameters | FLOAT
206 | ParM62 (Rel.density (15)) | Request -Method parameters | FLOAT
208 | ParM64 (Compression factor (15)) | Request -Method parameters | FLOAT
210 | ParM34 (Q Sup.vol.(0,25), MJ/m3) | Request -Method parameters | FLOAT
212 | ParM37 (Q Inf.vol.(0,25), MJ/m3) | Request -Method parameters | FLOAT
214 | ParM40 (W Sup.(0,25), MJ/m3) | Request -Method parameters | FLOAT
216 | ParM43 (W Inf.(0,25), MJ/m3) | Request -Method parameters | FLOAT
218 | ParM50 (Q Sup.vol.(0,25),Kwh/m3) | Request -Method parameters | FLOAT
220 | ParM47 (Q Inf.vol.(0,25),Kwh/m3) | Request -Method parameters | FLOAT
222 | ParM53 (W Sup.(0,25), Kwh/m3) | Request -Method parameters | FLOAT
224 | ParM57 (W Inf.(0,25), Kwh/m3) | Request -Method parameters | FLOAT
226 | ParM61 (Abs.density(0), kgr/m3) | Request -Method parameters | FLOAT
228 | ParM63 (Rel.density (0)) | Request -Method parameters | FLOAT
230 | ParM65 (Compression factor (0)) | Request -Method parameters | FLOAT
232 | ParM5 (Q Sup.vol.(20,25), MJ/m3) | Request -Method parameters | FLOAT
234 | ParM4 (Q Inf.vol.(20,25), MJ/m3) | Request -Method parameters | FLOAT
236 | ParM10 (W Sup.(20,25), MJ/m3) | Request -Method parameters | FLOAT
238 | ParM11 (W Inf.(20,25), MJ/m3) | Request -Method parameters | FLOAT
240 | ParM45 (Q Inf.vol.(20,25),KWh/m3) | Request -Method parameters | FLOAT
242 | ParM55 (W Sup.(20,25), Kwh/m3) | Request -Method parameters | FLOAT
244 | ParM59 (W Inf.(20,25), Kwh/m3) | Request -Method parameters | FLOAT
246 | ParM8 (Abs.density(20), kgr/m3) | Request -Method parameters | FLOAT
248 | ParM9 (Rel.density (20)) | Request -Method parameters | FLOAT
250 | ParM2 (Compression factor (20)) | Request -Method parameters | FLOAT
252 | ParM7 (Q Sup.mol.(25) kJ/mole) | Request -Method parameters | FLOAT
254 | ParM6 (Q Inf.mol.(25) kJ/mole) | Request -Method parameters | FLOAT
256 | ParM14 (Q Sup.mass.(25), MJ/kgr) | Request -Method parameters | FLOAT
258 | ParM13 (Q Inf.mass.(25), MJ/kgr) | Request -Method parameters | FLOAT
260 | ParM3 (Gas molecular weight) | Request -Method parameters | FLOAT
262 | Conc (Test gas component concentration) | Database query parameters-Methane | FLOAT
264 | Conc (Test gas component concentration) | Database query parameters-Ethane | FLOAT
266 | Conc (Test gas component concentration) | Database query parameters-Propane | FLOAT
268 | Conc (Test gas component concentration) | Database query parameters-i-Butane | FLOAT
270 | Conc (Test gas component concentration) | Database query parameters-n-Butane | FLOAT
272 | Conc (Test gas component concentration) | Database query parameters-neo-Pentane | FLOAT
274 | Conc (Test gas component concentration) | Database query parameters-i-Pentane | FLOAT
276 | Conc (Test gas component concentration) | Database query parameters-n-Pentane | FLOAT
278 | Conc (Test gas component concentration) | Database query parameters-n-Hexane+ | FLOAT
280 | Conc (Test gas component concentration) | Database query parameters-Carbon dioxide | FLOAT
282 | Conc (Test gas component concentration) | Database query parameters-Nitrogen | FLOAT
284 | Conc (Test gas component concentration) | Database query parameters-Oxygen | FLOAT
286 | FlagStat (Flag sampling analyzes) | Database query parameters | UINT16
287 | KodBase (Database code) | Database query parameters | UINT16
288 | MonthBase (Sample Month) | Database query parameters | UINT16
289 | YearBase (Sample Year) | Database query parameters | UINT16
290 | DayBase (Sample Day) | Database query parameters | UINT16
291 | HourBase (Sample Hour) | Database query parameters | UINT16
292 | PozBase (Sample position in the analysis) | Database query parameters | UINT16
293 | Value 1(Water Dew-Point (C)) | Parameters of Analysis mode | FLOAT
295 | Value 2(Pipe Pressure (кPa)) | Parameters of Analysis mode | FLOAT
297 | Value (the numerical field 3) | Parameters of Analysis mode | FLOAT
299 | Value (the numerical field 4) | Parameters of Analysis mode | FLOAT
301 | Value (the numerical field 5) | Parameters of Analysis mode | FLOAT
303 | NumKal (Number of Test Gas Passport) | Database query parameters | UINT16
304 | DateCur (Current date) | System Settings | VARIANT_TIME
Number of parameters168 number of registers: 308
| Function Code | 16
Address: | Name | Type | Data type
286 | FlagStat (Flag sampling analyzes) | Database query parameters | UINT16
287 | KodBase (Database code) | Database query parameters | UINT16
288 | MonthBase (Sample Month) | Database query parameters | UINT16
289 | YearBase (Sample Year) | Database query parameters | UINT16
290 | DayBase (Sample Day) | Database query parameters | UINT16
291 | HourBase (Sample Hour) | Database query parameters | UINT16
292 | PozBase (Sample position in the analysis) | Database query parameters | UINT16
293 | Value 1(Water Dew-Point (C)) | Parameters of Analysis mode | FLOAT
295 | Value 2(Pipe Pressure (кPa)) | Parameters of Analysis mode | FLOAT
297 | Value (the numerical field 3) | Parameters of Analysis mode | FLOAT
299 | Value (the numerical field 4) | Parameters of Analysis mode | FLOAT
301 | Value (the numerical field 5) | Parameters of Analysis mode | FLOAT
303 | NumKal (Number of Test Gas Passport) | Database query parameters | UINT16
304 | DateSet (Date to set) | time Synchronization | VARIANT_TIME
Number of parameters14 number of registers: 22
Registers in adresses between 286-292 inclusively related to querying database, result of a query (averages n-hourly, n-daily, n-analyses) are placed into the registers 130-260 inclusively.
Examples:
Average daily 05.02.2006 00:00 to 06.02.2006 00:00:
FlagStat = 1
KodBase = 0
MonthBase = 2
YearBase = 2006
DayBase = 0
HourBase = 0
PosBase = 5
Average month 01.02.2006 00:00 to 01.03.2006 00:00:
FlagStat = 3
KodBase = 0
MonthBase = 2
YearBase = 2006
DayBase = 0
HourBase = 0
PosBase = 0
Average 2 hours from 15.02.2006 09:00 to 17.02.2006 11:00:
FlagStat = 10
KodBase = 0
MonthBase = 2
YearBase = 2006
DayBase = 15
HourBase = 9
PosBase = 50
Average 1 hour from 15.02.2006 09:00 to 15.02.2006 10:00
FlagStat = 10
KodBase = 0
MonthBase = 2
YearBase = 2006
DayBase = 15
HourBase = 9
PosBase = 1
The query is handled on writing PosBase.

## 8 Errors
FlagErr (Analysis errors)
Erorrs bitflag of the last analysis. It's not a current error state flag of the device, it's an error flag of the analysis results of which are placed into other registers.
Critical error | Bitflag error encodings
YES | bit 0: Missing connection with the device module
YES | bit 1: Temperature of the termozone exceeds the limit
YES | bit 2: Module temperature measurement malfunction
YES | bit 3: Module PID-regulator error
YES | bit 4: Module input power error
YES | bit 5: Module non-fatal error
YES | bit 6: Module fatal error
YES | bit 7: Detector module not ready
YES | bit 8: Periphery error (pressure of service gases, humidity, flowrate and etc.)
NO | bit 9: Carrier gas ballon pressure exceeds the limit
NO | bit 10: Test gas ballon pressure exceeds the limit
YES | bit 11: Lua script execution error triggered
YES | bit 12: Module ADC malfunction
NO | bit 13: Controlled parameter exceeds the warning limit
NO | bit 14: Controlled parameter exceeds the alarm limit
NO | bit 15: Flagerror flipped by the user
* every bit aggregates a family of possible module errors, every error is logged in the internal device journal and can be analyzed later with application software. This is only a flag that triggers the need for such analysis.
** bit 13 and bit 14, each calculated parameter can be limited with lower or upper bound, e.g. it's possible to set the limit on the component concentration or physical parameter, or even add synthetic calculated property and add a limit to it.
** bit 15 is legacy, couldn't be set in a process GC.